BiPAP® Harmony®

Please check with your local Respironics Representative to find out if this product is available for purchase in your country.

The BiPAP Harmony has been designed to provide efficient and comfortable non invasive ventilation for non dependant chronic respiratory insufficiency patients, such as sleep apnea patients, COPD, obese hypoventilation and some restrictive diseases.

Like the BiPAP Synchrony , the BiPAP Harmony provides Respironics advanced technology Digital Auto-TRAK Sensitivity, providing automatically optimized trigger setting even in case of changes in leaks and patient respiratory mechanics.

Vent Ramp and RiseTime also aid in the delivery of comfortable ventilation.

Compliance therapy and trends data collected from the Encore SmartCard and then managed by Encore Pro software help clinicians assess home ventilation effectiveness and adjust settings to maintain comfort and efficiency of therapy in the long term. Harmony is also compatible withStardust II polygraph to provide detailed information on patients ventilation on a breath-to-breath basis.

Traveling is easy with the Harmony, as it weighs only 1,8 kilograms and can be connected to an external battery with a 7 hrs autonomy at EPAP/IPAP of 4/20 cm H2O and 12 bmp.

Specifications

BiPAP® Harmony®

Ventilation Modes

CPAP, Spontaneous (S), Spontaneous/Timed (S/T)

Settings

| IPAP | 4 to 30 cmH2O |
| --- | --- |
| EPAP | 4 to 25 cmH2O |
| CPAP | 4 to 20 cmH2O |
| Breath Rate | 0 to 30 BPM |
| Inspiratory Time | 0.5 to 3.0 sec |
| Rise Time | 100 to 600 msec (1 to 6) |
| Vent Ramp | 0 to 45 min. |

Physical

| Dimensions | 24 cm (L) x 17 cm (W) x 11 cm (H) |
| --- | --- |
| Weight | 1,8 Kilograms |

Electrical

| AC Voltage | 100 to 240 V, 50/60 Hz |
| --- | --- |
| DC Voltage | 12 V |
| AC Current | 1.25 A Maximum |
| DC Current | 5.5 A Maximum |

Trigger and Cycle

| Triggering | Automatic triggers with Auto-Trak algorithm: - Shape Signal - Volume 6 cc above baseline |
| --- | --- |
| Cycling | Automatic cycling with Auto-TrakTM algorithm: - Spontaneous Expiratory Threshold (SET) - Shape Signal - Maximum Inspiratory Time of 3.0 sec - Flow Reversal |

Accessories

| External battery | 7 hrs autonomy at EPAP/IPAP of 4/20cmH2O and 12 bmp |
| --- | --- |
| Integrated humidifier | 5 levels of heating |
| Compliance software | Encore® Pro software and Encore SmartCard, can be connected to Stardust® II polygraph |